## Slide 1: Detection Of Subsurface Defects In Dental Enamel Using Multispectral Truncated-correlation Photothermal Coherence Tomography Imaging
Elnaz B.Shokouhi
Andreas Mandelis

## Slide 2
Outline
Introduction
Objectives
Methodology
Theory
Experiments and results
Conclusions and future works

## Slide 3
Enhanced-Truncated Correlation Photothermal Coherence Tomography (eTC-PCT) methodology
Can harness photothermal phenomena to meet the demands for high-resolution 3D imaging
Overcomes poor axial resolution of thermophotonic technologies through time-evolving axial filtering controlled by pulse delay and cross-correlation truncation
SNR measurements for various chirp waveforms: sine, square, and fixed pulse-width chirps [1]
t
Sine chirp
Square chirp
Fixed pulsewidth
1. Kaiplavil, S., & Mandelis, A. Review of Scientific Instruments, 82(7), 074906 (2011) .

## Slide 4
eTC-PCT image reconstruction algorithm
1st reference chirp
2nd reference chirp
3rd reference chirp
Laser excitation chirp on sample
Synthesized reference chirps
Sample’s thermophotonic
relaxation chirp
Cross-correlate the signals pixel-by-pixel
Slice/2D images corresponding to the depth set by the delay applied to the reference chirp
Laser on

## Slide 5
Cross-correlation of relaxation chirp
and in-phase reference signal [2]
Cross-correlation of relaxation chirp
and quadrature reference signal [2]
Quadrature reference signal
In-phase Reference signal
Amplitude:
Phase:
Relaxation chirp for one pixel of hole 1
2. Tavakolian, P., Sivagurunathan, K., & Mandelis, A. Journal of Applied Physics, 122(2), 023103 (2017).
eTC-PCT image reconstruction: an example

## Slide 6
Background
Experimental parameters such as frequency, excitation pulse width, and wavelength effect depth penetration of eTC-PCT images
eTC-PCT image reconstruction was initially performed with a continuous wave (CW) 808 nm laser diode as the excitation source with a user-controlled pulse width and has been used in industrial and biomedical applications
Imaging application | Dental | Tumor | Brain
Depth resolution (mm) | ~3.8 [3] | ~2.0 [4] | NA [5]
CW laser excitation with controlled pulse width (ms) | 20-80 | 70 | 140
Biomedical applications of eTC-PCT to date
3. Roointanb, S., Tavakolian, P., Sivagurunathan, K. S., Mandelis, A., & Abrams, S. H. Journal of Biomedical Optics, 26(4), 046004 (2021).
4. Zhang, H., Tavakolian, P., Sivagurunathan, K., Mandelis, A., Shi, W., & Liu, F. F. Optics letters, 44(3), 675-678 (2019).
5. Tavakolian, P., Roointan, S., & Mandelis, A. Scientific reports, 10(1), 1-10 (2020).

## Slide 7
Optical sources such as pulsed Nd:YAG laser can generate ultra-short (nanosecond) pulses and an OPO (optical parametric oscillator) to develop multispectral TC-PCT
Multispectral TC-PCT allows us to choose the excitation wavelengths based on both the scattering and absorption coefficients
The amount of scattering is a function of the photon energy, that is, longer wavelengths are less scattered compared to shorter wavelengths (higher energy photons) can therefore penetrate deeper into the substructure
Objectives
End of pulse light
time
Tissue temperature
Excitation pulse duration vs.
thermal transient signal

## Slide 8
Methodology
Steps involved in multispectral TC-PCT system development:
Activating the second harmonic of an Nd:YAG laser
External trigger mode for added chirp pulse features
Methods for high energy laser pulse delivery (cross-polarizers, lens, beam expander, multimode fiber, and beam homogenizer)
Activating an OPO controlled software
Performing full pump beam alignment for OPO activation
OPO energy measurements
OPO internal system schematic

## Slide 9
Instrumental configuration
Multispectral TC-PCT
experimental setup
Q-switched Nd:YAG laser (Surelite OPO Plus SLIII-10, Continuum)
IR camera (A6700sc, FLIR, USA, 3–5 μm spectral response)
In multispectral TC-PCT, the laser excitation pulse width is fixed at 5 ns
Operates between 675-1000 nm wavelengths
Experiments are performed at single repetition frequency using the following equation:
= frequency of flashlamp discharge
= laser pulse repetition frequency

## Slide 10
Theory
The pulsed photothermal radiometry (PPTR) signal of a tissue subject to an instantaneous (nanosecond) pulse of Dirac delta type is represented by S(t):
The depth dependent optical subsurface distribution for a semi-infinite turbid sample is:

## Slide 11
Theory
Parameter study on theoretical signal profiles derived from S(t)
The effect of absorption
on the PPTR signal
The effect of scattering
on the PPTR signal
Parameters used for
soft tissues
- Absorption coefficient,
- Scattering coefficient,

## Slide 12
Sequence of experiments performed | Sample | Study target | Excitation Source Used
Experiment 1 |  | Comparative study between single repetition frequency vs. linear frequency modulation chirp of eTC-PCT images on a sample with artificially induced dental caries | CW laser
Experiment 2 |  | Comparative study between OPO wavelengths on a sample with artificially induced and natural dental caries | Pulsed Nd:YAG laser
Experiment 3 |  | Examining a natural dental crack on the basis of wavelength and single pulse repetition frequency | Pulsed Nd:YAG laser
Schematic imaging study sequence
Experimental Plan

## Slide 13
Comparison between eTC-PCT images of single repetition frequency and chirp frequency using CW laser
0.2 Hz
0.3 Hz
0.4 Hz
0.5 Hz
0.2-0.6 Hz
0.2 Hz
0.3 Hz
0.4 Hz
0.5 Hz
0.2-0.6 Hz
Artificial
caries
Amplitude
Phase
Before
demineralization
After
demineralization
1
0
Normalized Scale Bar
Cross-section
2 mm
2 mm
1 mm
Cross-section

## Slide 14
532 nm
675 nm
700 nm
750 nm
808 nm
L
1
0
Normalized Amplitude
532 nm
675 nm
700 nm
750 nm
808 nm
900 nm
z
x
2 mm
0.1 mm
900 nm
Amplitude at 0.5 Hz single repetition frequency
1 mm
1 mm
Top view
Multispectral TC-PCT imaging on dental caries

## Slide 15
532 nm
675 nm
750 nm
808 nm
700 nm
532 nm
675 nm
750 nm
808 nm
900 nm
700 nm
z
x
2 mm
0.1 mm
Phase at 0.5 Hz single repetition frequency
Top view
900 nm
Multispectral TC-PCT imaging on dental caries
1
0
Normalized phase
1 mm
1 mm

## Slide 16
Multispectral TC-PCT imaging on dental caries
Micro-CT slices of the tooth depicting the bacterial-induced caries and the caries captured with 532 nm OPO wavelength; and a deep natural feature in the tooth

## Slide 17
Multispectral TC-PCT imaging on dental cracks
Slice 1: 0 ms
Slice 2: 50 ms
Slice 3: 100 ms
Slice 4: 150 ms
Slice 5: 200 ms
Slice 6: 250 ms
Slice 7: 300 ms
Slice 8: 350 ms
Slice 9: 400 ms
Slice 10: 450 ms
Slice 1: 0 ms
Slice 2: 50 ms
Slice 3: 100 ms
Slice 4: 150 ms
Slice 5: 200 ms
Slice 6: 250 ms
Slice 7: 300 ms
Slice 8: 350 ms
Slice 9: 400 ms
Slice 10: 450 ms
Natural crack
1.2 cm
0.7 cm
Amplitude at 675 nm
Amplitude at 700 nm
Micro-CT slice of the crack

## Slide 18
Multispectral TC-PCT imaging on dental cracks
Natural crack
1.2 cm
0.7 cm
Phase at 675 nm
Slice 1: 0 ms
Slice 2: 50 ms
Slice 3: 100 ms
Slice 4: 150 ms
Slice 5: 200 ms
Slice 6: 250 ms
Slice 7: 300 ms
Slice 8: 350 ms
Slice 9: 400 ms
Slice 10: 450 ms
Slice 1: 0 ms
Slice 2: 50 ms
Slice 3: 100 ms
Slice 4: 150 ms
Slice 5: 200 ms
Slice 6: 250 ms
Slice 7: 300 ms
Slice 8: 350 ms
Slice 9: 400 ms
Slice 10: 450 ms
Phase at 700 nm

## Slide 19
Conclusions and future works
The successful preliminary multispectral TC-PCT imaging results for artificial and natural dental defects provided the rationale for further application of this technology in dental and bone imaging
With wide range of wavelengths and single repetition frequencies to choose from and the tissue-light scattering phenomena, we can detect near to the surface and deeper features
Multispectral TC-PCT is a good fit for future biomedical applications in imaging of small animals and human tissue imaging
Thank you!