## Slide 1
A. Mandelis1,2, A. Melnikov1,2, K. Sivagurunathan1,2
1 Center for advanced Diffusion-Wave and Photoacoustic Technologies (CADIPT), Department of Mechanical and Industrial Engineering, University of Toronto, Toronto, ON, M5S 3G8, Canada
2 Institute for Advanced Non-Destructive and Non-Invasive Diagnostic Technologies (IANDIT), Faculty of
Applied Science and Engineering, University of Toronto, ON, M5S 3G8, Canada
CADIPT
Quantitative non-destructive coating imaging and hardened steels case depth imaging using single frequency thermal-wave radar

## Slide 2
Experimental lock-in thermography (LIT) and thermal-wave radar (TWR) imaging setup
The imaging system consists of 5 main components:
  Infrared camera for thermal imaging
  Laser for excitation
  Automated 6-DOF (3D translational and 3D rotational axis) sample stage for easy access to region of interest (ROI)
  Monitoring camera for locating measuring site and autofocusing
  Computer and DAQ for system control and image processing
horizontal rotation
IR camera
monitoring camera
automated sample stage
computer
thermal emission
Z
Y
X
tilt
vertical rotation
laser
multi DAQ

## Slide 3
Conventional Lock-In Amplifier (LIA) signal processing.

## Slide 4
Conventional Digital Lock-In Thermography Imaging (LITI) signal processing.
n is determined by camera’s frame rate.
Increasing frequency leads to decreased n.
Synchronization with generating waveform
must be precise. Sync errors increase with frequency.
Undersampling at high frequencies creates additional stringent requirements for synchronization; LIT is not an optimal option for shallow-depth imaging.
Modulation frequency determines sampling length, so optimal frame rate is less than the maximum available
Co-addition (averaging) does not help with sync problems. Problem with low-cost cameras!
Result: a limited number of points is used for LIT amplitude and phase construction, and sync problems lead to low SNR especially at high frequencies.
Synchonous undersampling of a high frequency waveform using low sampling rate
Oversampling (below threshold frequency f0)
Undersampling (above threshold frequency f0)
Here: n is number of frames per correlation cycle;
Ij (x,y): value of x,y-pixel
sin(2πft): reference signal

## Slide 5
Review: Single-Frequency Thermal-Wave Radar Imaging (SF-TWRI) Signal Processing
Using the same frequency
for the start and end of a TWRI sweep allows the calculation
of the amplitude and phase at that frequency and the measurement of detailed pixel frequency dependencies by scanning over a wide range of frequencies at arbitrary intervals similar to photothermal radiometry, free from critical frequency exclusion, inter-frequency interval constraints, and upper limits imposed by sampling. Signal processing is done through CC, not through the lock-in process.
Critical parameters: single-frequency chirp duration + (max) camera frame rate ffr are operator controlled, not dictated by sampling theorem requirements.
Results: 1) The SNR is higher for SF-TWR than for LIT and the SNR difference increases with increasing frequency. 2) SF-TWR provides a uniform, continuous approach below and above the critical (undersampling boundary) frequency ffr /4.
SF-TWR is advantageous for work requiring frequency scans (quantitative dynamic thermography)
Chirp duration
Single frequency
TWRI sampling

## Slide 6: Comparison: LITI and SF-TWRI at 1 Hz for steel block with blind hole at 0.4 mm depth
Note horizontal dotted pixel lines

## Slide 7: Comparison: LITI and SF-TWRI at 1 Hz for steel block with blind hole at 0.4 mm depth (Cont’d)
Profile comparison between SF-TWRI and LITI amplitude and phase distributions at 16 Hz for the horizontal dotted pixel lines shown in previous slide

## Slide 8
Case I: Quantitative SF-TWRI: Two - layer theoretical model with grouped parameters for quantitative estimation of layer thickness.
Normalization by the semi-infinite
substrate signal
Where
, , ,
k0 = 0.026 Wm-1k-1
a0 = 22.26 x 10-6 m2/s
Parameters: Q1, P1, P2
deposited layer
air
substrate
ɑ0, k0
ɑ1, k1, L1
ɑ2, k2
Laser beam
The Q parameter is the most sensitive parameter to layer thickness and is most reliably measured through frequency-response best-fitting.

## Slide 9: Diagram of the SF-TWR imaging system

## Slide 10
Sample for imaging of coating thickness (Al substrate)
~40 x 40 x16 mm block
Al substrate, CoP nano-coating with nominal ~250 µm thickness
Y
X
25 mm
5.2 mm
Investigated area

## Slide 11
mm
0
25
20
5
10
15
TWR 2 Hz
TWR 10 Hz
TWR 20 Hz
SF-TWR phase images of sample with deposited Co-P alloy layer on the Al substrate.
A
B
Phase, degrees

## Slide 12: Phase. 8Hz. Deposited layer sample.
Normalization
to Zr sample
Fit to two-layer model
Phase. 8Hz. Zr
(phase of sample) – (phase of Zr)
Fitted Q - parameter
Phase normalization procedure

## Slide 13
Normalized phase dependence on frequency for ROI A and B
Phase was calculated as the mean value of a 6 x 6 pixel group in regions A and B.
The frequency of phase maximum fmax is determined by the parameter Q. Larger Q leads to fmax shift to lower frequency. Phase extrema appear due to standing thermal-wave across the coating layer thickness.

## Slide 14
Images of Q-parameter and thickness of deposited layer on the Al substrate.
mm
0
20
5
10
15
Thickness,
The diffusivity of the coated layer was assumed to be the same for all samples for calculating the coated layer thickness across the image. This assumption was proven to be generally valid.
The diffusivity was calculated from the best-fitted parameter Q at one point with independently measured thickness of the coated layer.
mm
0
20
5
10
15
C
C
0.1
0.05
0
200
140
80

## Slide 15
Directly measured and camera evaluated deposited layer thickness profiles
The apparent small divergence in regions close to the edges on both sides of the strip is likely due to a breakdown in the assumption of constant thermal diffusivity and/or one-dimensionality in those regions. Because the presence of the edges was not taken into account in the camera measurements, the thermal wave confinement would decrease the effective diffusivity within a thermal diffusion length from the edge, and the 1-D theory with constant diffusivity
predicted a smaller thickness as shown

## Slide 16: Case II: Quantitative SF-TWRI: case depth imaging of hardened steels
The imaged areas of two gear-tooth steel samples: (a) sample A11; (b) sample P6.

## Slide 17: Three - layer theoretical model with grouped parameters for quantitative estimation of hardened layer thickness.

## Slide 18: Case II. 1: Sample A11 hardness layer case depth NDI
Phase images of sample A11 with its own substrate as reference: (a) 1Hz; (b) 5Hz;
(c) 10Hz.

## Slide 19: Sample A11 hardness layer case depth NDI (Cont’d)

## Slide 20: Sample A11 hardness layer case depth NDI (Cont’d)

## Slide 21: Thermal diffusivity measurement of A11 hardened layer
Reference sample: Unhardened bulk of A11

## Slide 22: Sample A11 hardness layer case depth NDI (Cont’d)
Sample A11: (a) the fitted parameter image; (b) the constructed quantitative case depth image; (c) Vickers hardness profile. The average effective case depth of the highlighted area is 0.93 mm; the case depth according to the Vickers measurement is 1.05 mm.

## Slide 23: Case II. 2: Sample P6 hardness layer case depth NDI
Phase images of sample A11 with its own substrate as reference: (a) 1Hz; (b) 5Hz;
(c) 10Hz.

## Slide 24: Sample P6 hardness layer case depth NDI (Cont’d)

## Slide 25: Sample P6 hardness layer case depth NDI (Cont’d)

## Slide 26: Conclusions
Single-frequency (SF-) TWRI was used for quantitative NDI measurements of engineering solids with coatings or hardened case surface layers. The SNR is higher for SF-TWRI than for LITI and the SNR difference increases with increasing frequency.
Using camera frequency scans over extended frequency ranges not limited by the lock-in process allowed reconstructing quantitative thermophysical and geometric parameter images of materials.
Quantitative images of 1) a deposited Co-P alloy layer on Al substrate; 2) case depths of two hardened steels used in the aerospace industry, samples A11 and P6, were obtained using a two-layer and three-layer thermal-wave model, respectively.
The theory showed that the parameter Q = L/√a of the overlayer is the most sensitive and reliable parameter in fitting the frequency dependence of the mid-IR camera pixels to the thermal-wave model. This feature allowed the quantitative measurement of the coating and hardness layer thicknesses.

## Slide 27
University of Toronto
CADIPT
Acknowledgments
Natural Sciences and Engineering Research Council of Canada (NSERC) - Discovery grant
Canada Foundation for Innovation (CFI-JELF) - equipment grants
Canada Research Chairs Program